EDOC211

EUROGAM DATA ACQUISITION SYSTEM

Register Definitions for the FDT32 module

Edition 1.0

April 1995

Nuclear Physics Support Group

Central Laboratory of the Research Councils

Daresbury Laboratory

Registers for the fdt32 module

This document should be read in conjuction with the document EDOC210 which describes the Register Server implementation for CAMAC.

all register attributes include -c (crate address) and -n (station address)

Registers using the xCAMAC class of the CAMAC Register Server

control register		write by f(16) a(0)

The control register is 5 bits long and contains

bit 4	qerbit		value of data bit d30 on the DT32 bus - should 					normally be 0

bit 3	go_adc	enables data to be written by the ADCs to be 					FERA bus

bit 2	unused

bit 1	test		enables TEST mode - should normally be 0

bit 0	en_roco	enables the FERA bus to write to the fdt32 data  				FIFO and hence to the DT32 bus

fdt32#*.control		attributes  -a 0  -f 16  -w 16  -p wo  -l 0  -b 0  -i 0

fdt32#*.control.go_adc	attributes  -a 0  -f 16  -w 16  -p wo  -l 1  -b 3  -i 0

fdt32#*.control.test		attributes  -a 0  -f 16  -w 16  -p wo  -l 1  -b 1  -i 0

fdt32#*.control.en_roco	attributes  -a 0  -f 16  -w 16  -p wo  -l 1  -b 0  -i 0

status register 		read by f(0) a(0)

The status register is 9 bits long and contains

bit 8	en_roco	copy of control register bit 0

bit 7	dtpause	status of DTPAUSE signal on the DT32 bus

bit 6	dtstop		status of DTSTOP signal on the DT32 bus

bit 5	pass_fera	status of PASS signal from FERA bus

bit 4	dtren		status of DTREN signal from the DT32 bus

bit 3	notfull		status of FULL flag from FIFO (low active)

bit 2	notempty	status of EMPTY flag from FIFO (low active)

bit 1	nothfull	status of HALF FULL flag from FIFO (low active)

bit 0	dtpass		status of DTPASS signal on the DT32 bus

fdt32#*.status		attributes  -a 0  -f 0  -w 16  -p ro  -l 0  -b 0  -z x

fdt32#*.status.en_roco	attributes  -a 0  -f 0  -w 16  -p ro  -l 1  -b 8  -z d

fdt32#*.status.dtpause	attributes  -a 0  -f 0  -w 16  -p ro  -l 1  -b 7  -z d

fdt32#*.status.dtstop	attributes  -a 0  -f 0  -w 16  -p ro  -l 1  -b 6  -z d

fdt32#*.status.pass_fera	attributes  -a 0  -f 0  -w 16  -p ro  -l 1  -b 5  -z d

fdt32#*.status.dtren		attributes  -a 0  -f 0  -w 16  -p ro  -l 1  -b 4  -z d

fdt32#*.status.notfull	attributes  -a 0  -f 0  -w 16  -p ro  -l 1  -b 3  -z d

fdt32#*.status.notempty	attributes  -a 0  -f 0  -w 16  -p ro  -l 1  -b 2  -z d

fdt32#*.status.nothfull	attributes  -a 0  -f 0  -w 16  -p ro  -l 1  -b 1  -z d

fdt32#*.status.dtpass	attributes  -a 0  -f 0  -w 16  -p ro  -l 1  -b 0  -z d

write  gate delay	 DAC		write by f(16) a(6)

time between VALIDATION and start of GATE signal to the ADCs

adjustable between 120ns and 23us

fdt32#*.gate.delay	attributes  -a 6  -f 16  -w 16  -p wo  -l 0  -b 0

a sensible initial value is to be determined

write  gate length DAC	write by f(16) a(7)

length of GATE signal to the ADCs

adjustable between 50ns and 7us

fdt32#*.gate.length	attributes  -a 7  -f 16  -w 16  -p wo  -l 0  -b 0

a sensible initial value is to be determined

write  FT to validation delay DAC		write by f(16) a(8)

time between Fast Trigger and Validation signal

adjustable between 130ns and 20us

fdt32#*.val.delay	attributes  -a 8  -f 16  -w 16  -p wo  -l 0  -b 0

a sensible initial value is to be determined

write request to REN delay	 DAC		write by f(16) a(9)

time between REQUEST and start of ADC readout

adjustable between 75ns and 11us

fdt32#*.ren.delay	attributes  -a 9  -f 16  -w 16  -p wo  -l 0  -b 0

a sensible initial value is to be determined

select test output # 1		write by f(17) a(0)

various internal signals can be routed for inspection to the TEST OUT 1 lemo socket

data	0		gate

1		fast trigger (FT)

2		FT to validation delay

3		REQ

4		validation

5		CLI

6		trigger OK

7		WAK

15		WST from ADC

23		PASS from ADC

31		inhibit request

fdt32#*.test.#1	attributes  -a 0  -f 17  -w 16  -p wo  -l 0  -b 0

select test output # 2		write by f(17) a(1)

various internal signals can be routed for inspection to the TEST OUT 2 lemo socket

data	1		validok

2		endsubevent

3		/OPBM

4		/OPMBR

5		dtpause

6		dtren

7		dtpass

fdt32#*.test.#2		attributes  -a 1  -f 17  -w 16  -p wo  -l 3  -b 0

Note - this is a 3 bit register occupying the least significant bits of the data word

select test output # 3		write by f(17) a(1)

various internal signals can be routed for inspection to the TEST OUT 3 lemo socket

data	1		/OE

2		/WE

3		/RD

4		/WR

5		/HF

6		/FULL

7		/EMPTY

fdt32#*.test.#3		attributes  -a 1  -f 17  -w 16  -p wo  -l 3 -b 3

Note this is a 3 bit register which shares CAMAC access with the select test output # 2 function

Write LUT item		write by f(16) a(1)

Write a single data word to the fdt32 Look Up Table.

fdt32#*.writelutitem		attributes  -a 1  -f 16  -w 16  -p wo  -l 0 -b 0

Read LUT item		read by f(0) a(1)

Read a single data word from the fdt32 Look Up Table.

fdt32#*.readlutitem		attributes  -a 1  -f 0  -w 16  -p ro  -l 0 -b 0 -z x

Read FIFO item		read by f(0) a(2)

Read a single data word from the fdt32 data FIFO.

fdt32#*.readfifoitem		attributes  -a 2  -f 0  -w 24  -p ro  -l 0 -b 0 -z x

Registers using the cCAMAC class of the CAMAC Register Server

These registers do not require data and have no initial value. However read-register operations may return Q and X as data.

reset				write or read by f(9) a(0)

module reset

fdt32#*.reset		attributes  -a 0  -f 9

clear address counter	write or read by f(10) a(0)

set address used for access to the LUT to zero

fdt32#*.clearaddr		attributes  -a 0  -f 10

clear FIFO			write or read by f(9) a(2)

clear data FIFO

fdt32#*.clearfifo		attributes  -a 2  -f 9

enable LAM		write or read by f(26) a(0)

enable module LAM

fdt32#*.enablelam		attributes  -a 0  -f 26

disable LAM		write or read by f(24) a(0)

disable module LAM

fdt32#*.disablelam		attributes  -a 0  -f 24

test LAM			read by f(8) a(0)

returns Q=1 if the module LAM is set

fdt32#*.testlam		attributes  -a 0  -f 8

Registers using the qCAMAC class of the CAMAC Register Server

These registers transfer data using Q repeat mode between a file supplied by the write-register command and the fdt32 module. The write-register (or initialise-register) command is used to trigger execution of the CAMAC data transfer. The read-register command does not cause any CAMAC data transfer but just returns the file name supplied by the last write-register (or initialise-register) command.

Write LUT			write by f(16) a(1)

The fdt32 Look Up Table is 8K words of 16 bits.

fdt32#*.writelut		attributes  -a 1  -f 16  -w 16  -p wo  -l 8192

-i /eg/experiments/fdt32/LUT.default

Read LUT			read by f(0) a(1)

The fdt32 Look Up Table is 8K words of 16 bits.

fdt32#*.readlut		attributes  -a 1  -f 0  -w 16  -p ro  -l 8192

-i /eg/experiments/fdt32/LUT.current

Read FIFO			read by f(0) a(2)

The fdt32 data FIFO is 4K words of 24 bits.

fdt32#*.readfifo		attributes  -a 2  -f 0  -w 24  -p ro  -l 4096

-i /eg/experiments/fdt32/FIFO.current