EDOC213

CDOC213

EUROGAM DATA ACQUISITION SYSTEM

CHARISSA DATA ACQUISITION SYSTEM

Register Definitions for the Silena 4418/V ADC

Edition 1.0

April 1995

Nuclear Physics Support Group

Central Laboratory of the Research Councils

Daresbury Laboratory

Registers for the Silena 4418/V ADC module

This document should be read in conjunction with the document EDOC210 which describes the Register Server implementation for CAMAC.

This is a 8 channel ADC module with FERA and CAMAC data readout.

All register attributes include -c (crate address) and -n (station address)

Registers using the xCAMAC class of the CAMAC Register Server

status register 		read by f(4) a(14)  &  write by f(20) a(14)

The status register is composed of two distinct sections. The least significant 8 bits contain the Virtual Station Number (VSN) which determines the address of the data source during FERA readout. The most significant 8 bits determine the readout modes.

bit 8	unused	=0

bit 9	SUB		Channel Subaddress Enable, enabled with 					SUB=0. channel address is inserted into bits 12-14 				of the 	16 bit ADC data word.

bit 10	EEN		ECL port Enable: when EEN=1 ECL port readout 				is permitted, when EEN=0 CAMAC readout is 					permitted.

bit 11	OVF		Overflow Indication Enable, enabled with OVF=0.

overflow indication is inserted into bit 15 of the 16

bit ADC data word

bit 12	CCE		CAMAC data Compression Enable: when CCE=1, 				data zeros or data zeros-and-overflows are 					suppressed for CAMAC sequential readout

bit 13	CSR		CAMAC Sequential Readout: when CSR=0, 					CAMAC random access readout is enabled; when 				CSR=1, CAMAC sequential readout is enabled

bit 14	CLE		CAMAC Lam Enable: when CLE=1, LAM is 					enabled

bit 15	unused	=0

s4418#*.status	attributes  -a 14  -f 4  -w 16  -p rw  -l 0  -b 0  -z x -i 0

s4418#*.status.vsn		attributes  -a 14  -f 4  -w 16  -p rw  -l 8  -b 0  -z x

s4418#*.status.mode	attributes  -a 14  -f 4  -w 16  -p rw  -l 8  -b 8  -z x

s4418#*.status.sub		attributes  -a 14  -f 4  -w 16  -p rw  -l 1  -b 9  -z d

s4418#*.status.een		attributes  -a 14  -f 4  -w 16  -p rw  -l 1  -b 10  -z d

s4418#*.status.ovf		attributes  -a 14  -f 4  -w 16  -p rw  -l 1  -b 11  -z d

s4418#*.status.cce		attributes  -a 14  -f 4  -w 16  -p rw  -l 1  -b 12  -z d

s4418#*.status.csr		attributes  -a 14  -f 4  -w 16  -p rw  -l 1  -b 13  -z d

s4418#*.status.cle		attributes  -a 14  -f 4  -w 16  -p rw  -l 1  -b 14  -z d

Upper Level Discriminator	read by f(1) a(x) & write by f(17) a(x)

The upper threshold memory (8 bits) is accessed for channel x (0-7) using subaddress x. Values 0-255 correspond to 85% to 100% of full scale.

s4418#*.adc*.uld	attributes  -a x  -f 1  -w 16  -p rw  -l 8  -b 0  -i 255

Initial value is the current default used by the Charissa DA system.

Lower Level Discriminator	read by f(1) a(8+x) & write by f(17) a(8+x)

The lower threshold memory (8 bits) is accessed for channel x (0-7) using subaddress 8+x. Values 0-255 correspond to 0% to 10% of full scale.

s4418#*.adc*.lld	attributes  -a 8+x  -f 1  -w 16  -p rw  -l 8  -b 0  -i 1

Initial value is the current default used by the Charissa DA system.

Common Threshold	read by f(4) a(9) & write by f(20) a(9)

The common threshold memory (8 bits) applies to all channels.

s4418#*.threshold		attributes  -a 9  -f 4  -w 16  -p rw  -l 8  -b 0  -i 28

Initial value is the current default used by the Charissa DA system.

Offset Memory		read by f(4) a(x) & write by f(20) a(x)

The offset memory (8 bits) is accessed for channel x (0-7) using subaddress x.  Values 0-255 correspond to -3% to +3% of full scale. Value 128 is offset 0.

s4418#*.adc*.offset		attributes  -a x  -f 4  -w 16  -p rw  -l 8  -b 0  -i 128

Initial value is the current default used by the Charissa DA system.

ADC data 			read by f(0) a(x)  or  f(2) a(x)

Random access or sequential readout of the 8 ADC values via the CAMAC dataway. Channel x (0-7) is selected using subaddress x in random access readout only (CSR=0). In sequential mode (CSR=1) the subaddress is ignored.

Readout using f(2) clears the module.

s4418#*.adc*.data		attributes  -a x  -f 0  -w 16  -p ro  -l 0  -b 0

s4418#*.adc*.data+clr	attributes  -a x  -f 2  -w 16  -p ro  -l 0  -b 0

Header word data	read by f(0) a(14)  or  f(2) a(14)

Reads the event header word data f(0) a(14) or reads the event header word data and clears the module f(2) a(14).

s4418#*.hdr.data		attributes  -a 14  -f 0  -w 16  -p ro  -l 0  -b 0

s4418#*.hdr.data+clr	attributes  -a 14  -f 2  -w 16  -p ro  -l 0  -b 0

Pattern word data	read by f(0) a(15)  or  f(2) a(15)

Reads the event pattern word data f(0) a(15) or reads the event pattern word data and clears the module f(2) a(15).

s4418#*.patt.data		attributes  -a 15  -f 0  -w 16  -p ro  -l 0  -b 0

s4418#*.patt.data+clr	attributes  -a 15  -f 2  -w 16  -p ro  -l 0  -b 0

Registers using the cCAMAC class of the CAMAC Register Server

These registers do not require data and have no initial value. However read-register operations may return Q and X as data.

reset				write or read by f(9) a(0)

module reset

s4418#*.reset		attributes  -a 0  -f 9

test LAM			read by f(8) a(0)

returns Q=1 if the module LAM is set

s4418#*.testlam		attributes  -a 0  -f 8

test LAM and clear	read by f(10) a(0)

returns Q=1 if the module LAM is set and clears the LAM

s4418#*.testclearlam	attributes  -a 0  -f 10

test module		read by f(25) a(0)

returns Q=1 if BUSY =0,  returns Q=0 if the module is not in the “ready” state.

s4418#*.test			attributes  -a 0  -f 25