CDOC218

CHARISSA DATA ACQUISITION SYSTEM

Register Definitions for the LeCroy 2551 Scaler

Edition 1.0

May 1995

Nuclear Physics Support Group

Central Laboratory of the Research Councils

Daresbury Laboratory

Registers for the LeCroy 2551 Scaler module

This document should be read in conjunction with the document EDOC210 which describes the Register Server implementation for CAMAC.

This is a 12 channel Scaler module with CAMAC data readout.

All register attributes include -c (crate address) and -n (station address)

Registers using the xCAMAC class of the CAMAC Register Server

Channel data 			read by f(0) a(x)  or  f(2) a(x)

Addressed readout of the 12 register values (channels) via the CAMAC dataway. Channel x (0-11) is selected using subaddress x.

Readout using f(2) clears the module and LAM on A(11).

lc2551#*.c[0-11].data	      attributes  -a x  -f 0  -w 16  -p ro  -l 0  -b 0  -z x

lc2551#*.c[0-11].data+clr	      attributes  -a x  -f 2  -w 16  -p ro  -l 0  -b 0  -z x

Registers using the cCAMAC class of the CAMAC Register Server

These registers do not require data and have no initial value. However read-register operations may return Q and X as data.

reset				write or read by f(9) a(0)

Module reset. Clears all scaler channels simultaneously.

lc2551#*.reset		attributes  -a 0  -f 9

test LAM			read by f(8) a(0)

returns Q=1 if the module LAM is set

lc2551#*.testlam		attributes  -a 0  -f 8

Disable LAM		read by f(24) a(0)

Disables the module LAM. The state of the module LAM enable will be arbitary after power on.

lc2551#*.disable		attributes  -a 0  -f 24

Enable LAM		read by f(26) a(0)

Enables the module LAM.

lc2551#*.enable		attributes  -a 0  -f 26

Increment scalers		read by f(25) a(0)

Increment all scalers.

lc2551#*.increment		attributes  -a 0  -f 25