CDOC219

CHARISSA DATA ACQUISITION SYSTEM

Register Definitions for the LeCroy 2277 TDC

Edition 1.0

May 1995

Nuclear Physics Support Group

Central Laboratory of the Research Councils

Daresbury Laboratory

Registers for the LeCroy 2277 TDC module

This document should be read in conjunction with the document EDOC210 which describes the Register Server implementation for CAMAC.

This is a 32 channel TDC module with CAMAC data readout.

All register attributes include -c (crate address) and -n (station address)

Registers using the xCAMAC class of the CAMAC Register Server

status register 		read by f(1) a(0) and write by f(17) a(0)

The data value stored in the status register is interpreted by the module as a bit pattern.

bits 0,1	Tester hit burst		00 = 1 hit

01 = 2 hits

10 = 4 hits

11 = 8 hits

bits 2,3	Tester hit duration		00 = 128 nsec

01 = 256 nsec

10 = 1024 nsec

11 = 2048 nsec

bits 4,5,6	unused			=0

bit 7		Trigger mode		0 = pulse:	1 = latch

bit 8		odd channel test input	0 = disable	1 = enable

bit 9		even channel test input	0 = disable	1 = enable

bit 10		hit trailing edge recording	0 = disable	1 = enable

bit 11		hit leading edge recording	0 = disable	1 = enable

bits 12-14	common start timeout	000 = front panel

001 = 1024 nsec

010 = 2048 nsec

011 = 4096 nsec

100 = 8192 nsec

101 = 16384 nsec

110 = 32768 nsec

111 = 65536 nsec

bit 15		mode			0 = common stop	1=common start

lc2277#*.status		attributes  -a 0  -f 1  -w 16  -p rw  -l 0  -b 0  -z x -i 0

lc2277#*.status.mode      attributes  -a 0  -f 1  -w 16  -p rw  -l 1  -b 15  -z d -i 1

lc2277#*.status.timeout   attributes  -a 0  -f 1  -w 16  -p rw  -l 3  -b 12  -z d -i 5

lc2277#*.status.leading	attributes  -a 0  -f 1  -w 16  -p rw  -l 1  -b 11  -z d

lc2277#*.status.trailing    attributes  -a 0  -f 1  -w 16  -p rw  -l 1  -b 10  -z d -i 1

lc2277#*.status.eventest	attributes  -a 0  -f 1  -w 16  -p rw  -l 1  -b 9  -z d

lc2277#*.status.oddtest	attributes  -a 0  -f 1  -w 16  -p rw  -l 1  -b 8  -z d

lc2277#*.status.trigmode	attributes  -a 0  -f 1  -w 16  -p rw  -l 1  -b 7  -z d

lc2277#*.status.testhit	attributes  -a 0  -f 1  -w 16  -p rw  -l 2  -b 2  -z d

lc2277#*.status.testhitlen	attributes  -a 0  -f 1  -w 16  -p rw  -l 2  -b 0  -z d

Channel data 			read by f(0) a(0)

Data readout via the CAMAC dataway. Module returns Q=1 if valid data is present; Q=0 if FIFO is empty.

lc2277#*.data		attributes  -a 0 -f 0  -w 24  -p ro  -l 0  -b 0  -z x  -q 1

Registers using the cCAMAC class of the CAMAC Register Server

These registers do not require data and have no initial value. However read-register operations may return Q and X as data.

Clear				write or read by f(9) a(0)

Module clear. F(9) does not affect the status register.

lc2277#*.clear		attributes  -a 0  -f 9

Test LAM			read by f(8)  a(0)

returns Q=1 if the module LAM is set

lc2277#*.testlam		attributes  -a 0  -f 8

Clear LAM			read by f(10)  a(0)

clears the module LAM

lc2277#*.clearlam		attributes  -a 0  -f 10

Disable LAM		read by f(24) a(0)

Disables the module LAM. The state of the module LAM enable will be arbitary after power on.

lc2277#*.disable		attributes  -a 0  -f 24

Enable LAM		read by f(26) a(0)

Enables the module LAM.

lc2277#*.enable		attributes  -a 0  -f 26

Test cycle			read by f(25) a(0)

Test cycle.

lc2277#*.test			attributes  -a 0  -f 25

Test “buffering in progress”	read by f(27) a(0)

returns Q=1 if true

lc2277#*.testbuffering	attributes  -a 0  -f 27

Test Busy			read by f(27) a(1)

returns Q=1 if Busy is true.

lc2277#*.testbusy		attributes  -a 1  -f 27