EDOC221

CDOC221

EUROGAM DATA ACQUISITION SYSTEM

CHARISSA DATA ACQUISITION SYSTEM

Register Definitions for the GANELEC QDC1612F  ADC

Edition 1.0

Oct 1995

Nuclear Physics Support Group

Central Laboratory of the Research Councils

Daresbury Laboratory

Registers for the GANELEC QDC1612F ADC module

This document should be read in conjunction with the document EDOC210 which describes the Register Server implementation for CAMAC.

This is a 8 channel dual integration ADC module with FERA and CAMAC data readout.

All register attributes include -c (crate address) and -n (station address)

Registers using the xCAMAC class of the CAMAC Register Server

status register 		read by f(0) a(0)  &  write by f(16) a(0)

The status register controls the readout mode and allows the duration of the Master gate to be selected.

Note - bit numbering is bit 0 (LSB) and bit 15 (MSB).

bit 0	EPS		ECL port Pedestal Subtraction: when EPS=1,					pedestals are subtracted for ECL port readout

bit 1	ECE		ECL port data Compression Enable: when ECE=1,				data zeros or zeros-and-overflows are suppressed 				for ECL port readout.

bit 2	EEN		ECL port Enable: when EEN=1 ECL port readout 				is permitted, when EEN=0 CAMAC readout is 					permitted.

bit 3	CPS		CAMAC Pedestal Subtraction: when CPS=1,					pedestals are subtracted for CAMAC readout

bit 4	CCE		CAMAC data Compression Enable: when CCE=1,				data zeros or zeros-and-overflows are suppressed 				for CAMAC sequential readout.

bit 5	CSR		CAMAC Sequential Readout: when CSR=0, 					CAMAC random access readout is enabled; when 				CSR=1, CAMAC sequential readout is enabled

bit 6	CLE		CAMAC Lam Enable: when CLE=1, LAM is 					enabled

bit 7	OFS		OverFlow Suppression: when OFS=1, data					readouts are suppressed during ECL port readout				in conjunction with ECE=1 and for CAMAC					sequential readout in conjunction with CCE=1

bit 8	INDIV		INDIVidual. If INDIV=1, only the individual gates				FIGi and SIGi are authorised; if INDIV=0, only the 				common gates FCG and SCG are enabled.

bits 9 to 12 (4 bits)	Duration of the Master gate in microsecs.					Maximum duration is equal to 16 microsecs.

qdc1612#*.status		attributes  -a 0  -f 0  -w 16  -p rw  -l 0  -b 0  -z x -i 0

qdc1612#*.status.mode	attributes  -a 0  -f 0  -w 16  -p rw  -l 9  -b 0  -z x

qdc1612#*.status.Mgate	attributes  -a 0  -f 0  -w 16  -p rw  -l 4  -b 9  -z d

qdc1612#*.status.eps	attributes  -a 0  -f 0  -w 16  -p rw  -l 1  -b 0  -z d

qdc1612#*.status.ece	attributes  -a 0  -f 0  -w 16  -p rw  -l 1  -b 1  -z d

qdc1612#*.status.een	attributes  -a 0  -f 0  -w 16  -p rw  -l 1  -b 2  -z d

qdc1612#*.status.cps	attributes  -a 0  -f 0  -w 16  -p rw  -l 1  -b 3  -z d

qdc1612#*.status.cce	attributes  -a 0  -f 0  -w 16  -p rw  -l 1  -b 4  -z d

qdc1612#*.status.csr	attributes  -a 0  -f 0  -w 16  -p rw  -l 1  -b 5  -z d

qdc1612#*.status.cle	attributes  -a 0  -f 0  -w 16  -p rw  -l 1  -b 6  -z d

qdc1612#*.status.ofs	attributes  -a 0  -f 0  -w 16  -p rw  -l 1  -b 7  -z d

qdc1612#*.status.indiv	attributes  -a 0  -f 0  -w 16  -p rw  -l 1  -b 8  -z d

Virtual Station Number #1	read by f(0) a(1) & write by f(16) a(1)

Accesses the virtual station number VSN1 of the first group of analog inputs 1 to 4 (8 bits).

qdc1612#*.vsn1	attributes  -a 1  -f 0  -w 16  -p rw  -l 8  -b 0 -z d

Virtual Station Number #2	read by f(0) a(2) & write by f(16) a(2)

Accesses the virtual station number VSN2 of the second group of analog inputs 5 to 8 (8 bits).

qdc1612#*.vsn2	attributes  -a 2  -f 0  -w 16  -p rw  -l 8  -b 0 -z d

Pedestal memory	read by f(1) a(x) and write by f(17) a(x)

The pedestal memory (8 bits) is accessed for integrating channel x (0-15) using subaddress x. Integrating channel 1 corresponds to analog input 1 and narrow integrating gate; integrating channel 2 corresponds to analog input 1 and wide integrating gate etc.

qdc1612#*.channel*.pedestal	attributes  -a x  -f 1  -w 16  -p rw  -l 8  -b 0

ADC data 			read by f(2) a(x)

random access or sequential readout of the 16 16 bit data words via the CAMAC dataway. Channel x (0-15) is selected using subaddress x in random access readout only (CSR=0). In sequential mode (CSR=1) the subaddress is ignored.

qdc1612#*.channel*.data	attributes  -a x  -f 2  -w 16  -p ro  -l 0  -b 0

DAC Voltage		write by f(16) a(3)

Writes the voltage for the DAC delivering a common pedestal current to all integrating channels (8 bits).

qdc1612#*.Vdac		attributes  -a 3  -f 16  -w 16  -p wo  -l 8  -b 0

Registers using the cCAMAC class of the CAMAC Register Server

These registers do not require data and have no initial value. However read-register operations may return Q and X as data.

reset				write or read by f(9) a(0)

module reset

qdc1612#*.reset		attributes  -a 0  -f 9

test LAM			read by f(8) a(0)

returns Q=1 if the module LAM is set

qdc1612#*.testlam		attributes  -a 0  -f 8

test LAM and clear	read by f(10) a(0)

returns Q=1 if the module LAM is set and clears the LAM

qdc1612#*.testclearlam	attributes  -a 0  -f 10