EDOC226

Mi d a s

Register Definitions for the Ortec 413 ADC

Edition 1.0

Dec 1995

Nuclear Physics Support Group

Central Laboratory of the Research Councils

Daresbury Laboratory

Registers for the Ortec 413 ADC module

This document should be read in conjunction with the document EDOC210 which describes the Register Server implementation for CAMAC.

This is a 4 channel ADC module with FERA and CAMAC data readout.

All register attributes include -c (crate address) and -n (station address)

Registers using the xCAMAC class of the CAMAC Register Server

control register 1	read by f(0) a(0)  &  write by f(16) a(0)

This control register is composed of two distinct sections. The least significant 8 bits contain the Virtual Station Number (VSN) which determines the address of the data source during FERA readout. The most significant 8 bits determine the readout modes.

Note that most of the control bits use bit=0 to enable the option and hence the register enables the option when value=0.

bit 8	ZSE		Zero-Suppression Enable. Specifies zero-					suppression readout (ZSE=0) as opposed to					sequential mode (ECL or CAMAC) or random					access (CAMAC) (ZSE=1) (see CSR).

bit 9	EEN		ECL port Enable (EEN=0) or CAMAC readout					(EEN=1)

bit 10			not used

bit 11			not used

bit 12	CME		Specifies coincidence mode (CME=0) or singles				mode (CME=1).

bit 13	CSR		Specifies CAMAC Sequential Readout (CSR=0)				or CAMAC random access readout (CSR=1).

bit 14	CLE		Asserts CAMAC LAM (CLE=1) when data is ready				to be readout on the CAMAC port. If CLE=0 no					LAM is asserted.

bit 15	OFS		OverFlow Suppression: when OFS=0 then data is				suppressed (ZSE=0) or reported as zero (ZSE=1).

If OFS=1 then data overflows are reported with a				value between 8064 and 8191.

ot413#*.control1		attributes  -a 0  -f 0  -w 16  -p rw  -l 0  -b 0  -z x -i 0

ot413#*.control.vsn		attributes  -a 0  -f 0  -w 16  -p rw  -l 8  -b 0  -z x

ot413#*.control.mode	attributes  -a 0  -f 0  -w 16  -p rw  -l 8  -b 8  -z x

ot413#*.control.zse		attributes  -a 0  -f 0  -w 16  -p rw  -l 1  -b 8  -z d

ot413#*.control.een		attributes  -a 0  -f 0  -w 16  -p rw  -l 1  -b 9  -z d

ot413#*.control.cme		attributes  -a 0  -f 0  -w 16  -p rw  -l 1  -b 12  -z d

ot413#*.control.csr		attributes  -a 0  -f 0  -w 16  -p rw  -l 1  -b 13  -z d

ot413#*.control.cle		attributes  -a 0  -f 0  -w 16  -p rw  -l 1  -b 14  -z d

ot413#*.control.ofs		attributes  -a 0  -f 0  -w 16  -p rw  -l 1  -b 15  -z d

control register 2	read by f(0) a(1)  &  write by f(16) a(1)

This control register contains 5 bits which determine which gate signals are enabled. If a gate signal is disabled then a pulse is accepted regardless of the state of the gate signal.

Again note that the register enables the option when value=0.

bit 0	gate1		Front-panel gate for channel 1 enabled (gate1=0) - 				gate for channel 1 disabled (gate1=1).

bit 1	gate2		Front-panel gate for channel 2 enabled (gate2=0) - 				gate for channel 2 disabled (gate2=1).

bit 2	gate3		Front-panel gate for channel 3 enabled (gate3=0) - 				gate for channel 3 disabled (gate3=1).

bit 3	gate4		Front-panel gate for channel 4 enabled (gate4=0) - 				gate for channel 4 disabled (gate4=1).

bit 4	gatem		master gate enabled (mgate = 0) - master gate 				disabled (mgate = 1)

ot413#*.control2		attributes  -a 0  -f 1  -w 16  -p rw  -l 5  -b 0  -z x -i 0

ot413#*.gate1.enable	attributes  -a 0  -f 1  -w 16  -p rw  -l 1  -b 0  -z d

ot413#*.gate2.enable	attributes  -a 0  -f 1  -w 16  -p rw  -l 1  -b 1  -z d

ot413#*.gate3.enable	attributes  -a 0  -f 1  -w 16  -p rw  -l 1  -b 2  -z d

ot413#*.gate4.enable	attributes  -a 0  -f 1  -w 16  -p rw  -l 1  -b 3  -z d

ot413#*.gatem.enable	attributes  -a 0  -f 1  -w 16  -p rw  -l 1  -b 4  -z d

Lower level discrimator   read by f(1) a(x) and write by f(17) a(x)

The lower level discrimator (8 bits) is accessed for channel x (0-3) using subaddress x. The initial value (36) is the hardware default. The register value is in units of 2mV.

ot413#*.adc*.lld		attributes  -a x  -f 1  -w 16  -p rw  -l 8  -b 0 -i 36

ADC data 			read by f(2) a(x)

random access or sequential readout of the 4 ADC values via the CAMAC dataway. Channel x (0-3) is selected using subaddress x in random access readout only (CSR=1). In sequential mode (CSR=0) the subaddress is ignored. Also requires EEN=1 (CAMAC readout enabled) and ZSE=1.

ot413#*.adc*.data		attributes  -a x  -f 2  -w 16  -p ro  -l 0  -b 0

Registers using the cCAMAC class of the CAMAC Register Server

These registers do not require data and have no initial value. However read-register operations may return Q and X as data.

clear				write or read by f(9) a(0)

module clear

ot413#*.clear			attributes  -a 0  -f 9

test LAM			read by f(8) a(0)

returns Q=1 if the module LAM is set

lc4300b#*.testlam		attributes  -a 0  -f 8