DOC227

MIDAS

Register Definitions for the LeCroy 4434 Scaler

Edition 1.0

Jan 1996

Nuclear Physics Support Group

Central Laboratory of the Research Councils

Daresbury Laboratory

Registers for the LeCroy 4434 Scaler module

This document should be read in conjunction with the document EDOC210 which describes the Register Server implementation for CAMAC.

This is a 32 Channel Latching Scaler module with CAMAC data readout.

All register attributes include -c (crate address) and -n (station address)

Registers using the xCAMAC class of the CAMAC Register Server

Command Register			write by f(16) a(0)

The Command Register is a 16-bit word which controls the channel readout.

bits 0-4 (5 bits)	FA	First Address to be read.

bit 5			LD	Load - if LD=1 performs a load if switch

LAD (Latching Disable) = OFF.

bit 6			CL	Clear - if CL=1 clears all the 32 scalers

bit 7			RD	Readout Enable - if RD=1 prepares the

module for readout.

bits 8-12 (5 bits)	RN	Readout Number - defines how many

channels (minus 1) have to be read out

starting from the address FA.

bit 15			T	Test - if T=1 inhibits signal inputs and

increments all scalers.

lc4434#*.command			attributes  -a 0  -f 16  -w 16  -p wo  -l 0  -b 0

lc4434#*.command.FA		attributes  -a 0  -f 16  -w 16  -p wo  -l 5  -b 0

lc4434#*.command.RN		attributes  -a 0  -f 16  -w 16  -p wo  -l 5  -b 8

lc4434#*.command.T		attributes  -a 0  -f 16  -w 16  -p wo  -l 1 -b 15

lc4434#*.command.LD		attributes  -a 0  -f 16  -w 16  -p wo  -l 1  -b 5

lc4434#*.command.CL		attributes  -a 0  -f 16  -w 16  -p wo  -l 1  -b 6

lc4434#*.command.RD		attributes  -a 0  -f 16  -w 16  -p wo  -l 1  -b 7

Channel data 			read by f(0) a(0)

Generates readout of the selected register value (channel) via the CAMAC dataway.

lc4434#*.data		attributes  -a 0  -f 0  -w 16  -p ro  -l 0  -b 0  -z x

Channel data 			read by f(0) a(2)

Sequential data readout of the channel values with auto-increment of the address.

lc4434#*.data++		attributes  -a 2  -f 0  -w 16  -p ro  -l 0  -b 0  -z x

Registers using the cCAMAC class of the CAMAC Register Server

These registers do not require data and have no initial value. However read-register operations may return Q and X as data.

test LAM			read by f(8) a(0)

returns Q=1 if the module LAM is set

lc4434#*.testlam		attributes  -a 0  -f 8

test and clear LAM	read by f(10) a(0)

returns Q=1 if the module LAM is set and clears it

lc4434#*.test+clrlam	attributes  -a 0  -f 10