DOC231

MIDAS

Register Definitions for the LeCroy 4504 module

Edition 1.0

Jan 1996

Nuclear Physics Support Group

Central Laboratory of the Research Councils

Daresbury Laboratory

Registers for the LeCroy 4504 Flash ADC module

This document should be read in conjunction with the document EDOC210 which describes the Register Server implementation for CAMAC.

This is a four channel Flash ADC Module.

All register attributes include -c (crate address) and -n (station address)

Registers using the xCAMAC class of the CAMAC Register Server

Digitial Output		read by f(0) a(0)

Read digitial outputs; 4 bits per channel.

lc4504#*.adc0.data		attributes  -a 0  -f 0  -w 16  -p ro  -l 4  -b 0

lc4504#*.adc1.data		attributes  -a 0  -f 0  -w 16  -p ro  -l 4  -b 4

lc4504#*.adc2.data		attributes  -a 0  -f 0  -w 16  -p ro  -l 4  -b 8

lc4504#*.adc3.data		attributes  -a 0  -f 0  -w 16  -p ro  -l 4  -b 12

Digitial Overflow	read by f(1) a(0)

Read digitial overflows; 1 bit per channel.

lc4504#*.adc0.oflow	attributes  -a 0  -f 1  -w 16  -p ro  -l 1  -b 0 -z d

lc4504#*.adc1.oflow	attributes  -a 0  -f 1  -w 16  -p ro  -l 1  -b 1 -z d

lc4504#*.adc2.oflow	attributes  -a 0  -f 1  -w 16  -p ro  -l 1  -b 2 -z d

lc4504#*.adc3.oflow	attributes  -a 0  -f 1  -w 16  -p ro  -l 1  -b 3 -z d

Low Reference Voltage	write by f(16) a(0)

Write the Low Reference Voltage (VL), 8 bits; range from -2550 mV to +2250 mV in steps of 20 mV.

lc4504#*.VL			attributes  -a 0  -f 16  -w 16  -p wo  -l 8  -b 0

High Reference Voltage	write by f(16) a(1)

Write the High Reference Voltage (VH), 8 bits; range from -2550 mV to +2250 mV in steps of 20 mV.

lc4504#*.VH			attributes  -a 1  -f 16  -w 16  -p wo  -l 8  -b 0

Registers using the cCAMAC class of the CAMAC Register Server

These registers do not require data and have no initial value. However read-register operations may return Q and X as data.

C				write or read by f(25) a(0)

module C function

lc4504#*.C			attributes  -a 0  -f 25