DOC233

MIDAS

Register Definitions for the LeCroy 4532 module

Edition 1.0

Jan 1996

Nuclear Physics Support Group

Central Laboratory of the Research Councils

Daresbury Laboratory

Registers for the LeCroy 4532 module

This document should be read in conjunction with the document EDOC210 which describes the Register Server implementation for CAMAC.

This is a Majority Logic Unit.

All register attributes include -c (crate address) and -n (station address)

Registers using the xCAMAC class of the CAMAC Register Server

Input Pattern low		read by f(0) a(0)

Read the input pattern, channels 1 to 16.

lc4532#*.pattern.L		attributes  -a 0  -f 0  -w 16  -p ro  -l 0  -b 0

Input Pattern high		read by f(0) a(1)

Read the input pattern, channels 17 to 32.

lc4532#*.pattern.H		attributes  -a 1  -f 0  -w 16  -p ro  -l 0  -b 0

Status Register			read by f(1) a(0)

Read the status register.

bit 0	LAM			set = 1 if LAM is ON

bit 1	LAM enable		set = 1 if LAM Enable switch is ON

bit 2	MEMORY enable	set = 1 if MEMORY Enable switch is ON

bit 3	CLUSTER enable	set = 1 if CLUSTER Enable switch is ON

lc4532#*.status			attributes  -a 0  -f 1  -w 16  -p ro  -l 4  -b 0

lc4532#*.status.LAM		attributes  -a 0  -f 1  -p ro  -l 1  -b 0  -z d

lc4532#*.status.LAMenable	attributes  -a 0  -f 1  -p ro  -l 1  -b 1  -z d

lc4532#*.status.MEMenable	attributes  -a 0  -f 1  -p ro  -l 1  -b 2  -z d

lc4532#*.status.Cenable		attributes  -a 0  -f 1  -p ro  -l 1  -b 3  -z d

Registers using the cCAMAC class of the CAMAC Register Server

These registers do not require data and have no initial value. However read-register operations may return Q and X as data.

test lam 			write or read by f(8) a(0)

Test LAM; returns Q=1 if L is ON

lc4532#*.testlam		attributes  -a 0  -f 8

clear				read by f(9) a(0)

Clears the data memory and LAM

lc4532#*.clear		attributes  -a 0  -f 9

test and clear		read by f(10) a(0)

Test and Clear LAM; returns Q=1 if L is ON.

lc4532#*.test+clrlam	attributes  -a 0  -f 10